Manual Instruction for

Weighing machine  touchpanel – terminal

Only for the AKE 2-4  units

side		menu					description

3  		POWER ON				Power on menu

3-5		KEYWORD				password menu

6		DATA/TIME				change date and time

7		CONFIG				change windows settings

8		LANGUAGE				change language

9		BACKUP				backup and restore setting - files

10-11  		UPDATE SERVER			programme update

12		UPDATE CLIENTS  		! not for AKE-Units !

13     	INITIALISIERUNG		weighing machine initialization

14     	MACHINE				start weighing-programme

15     	COUNTER				set shedding-counter

16-17		OPERATING MODE		set/charge operating mode

18		RECIPES				database of recipes

19-20   	-  COPY				copy recipes

21-22  		-  EDIT				edit recipes

23    		-  DELETE				delete recipe

24    		-  SEARCHING			searching recipe with number

25    		-  SEARCH PRODUCTNAME 	searching recipe with name

26    		TARE SETTINGS			change tare-settings

27    		CHANGE DELAYS			change delay-times

28    		SUPPLY SETTINGS		change supply-settings

29    		REFERENCES SETTINGS	change reference values

30    		TARE					tare weighing head

31   		EMPTY 				empty weighing line or weighing bowl

32    		ON-OFF				weighing line switch ON or OFF

33    		ACTIVITY ADVERTISEMENT

34		CONTROL FUNCTIONS		weighing head, input, output controls

35 		WEIGH. HEAD CONTROL	weighing head control

36   		INPUT-OUTPUT CONTROL	input, output control

37		OPTIONS				setting optional functions

38-39 		TOUCHPANEL CALIBRATION	Calibration the touchpanel

40    		NOTIZE TO INITIALIZATION	read me !

41-42		NOTIZE TO OPTIMATION	description of Optimation

Menu-s

- Power-on menu

Only the button KEYWORD is active!

(  using another button , you will get an error  message )

Push the button KEYWORD to activate the password-menu:

Push  the accesslevel relevante button:

service = service, foreman = foreman,  operator = operator …

Please enter your 4-digit keyword. ( Every digit will be displayed by a star .)

If keyword-number is accepted,

you are asked whether to change your keyword or not.

Answer with NO if you don´t want to change your keyword

and you will leave the keyword-menu automatically.

If you answer with YES  the keyword-menu appears again.

Enter your new keyword into the upper window and repeat it in the lower window.

If the two keywords agree you will leave automatically,

if not you will get an „error“-message.

After confirming the message you have to repeat the procedure.

This message also appears if you enter a not registrated or wrong keyword.

Warranties:

SERVICE	free access to all functions

FOREMAN	no access 		-    to the initialization of the weighing machine						-    to update - menus

- to main-tare and adjustment

- to change or delete the „PROTECT“ recipes

OPERATOR	access free 		-     to start, stop, read recipes, simple- tara,  					       and modify references

2.   DATA/TIME  Enter or change date or time.

By pushing a window you can change the adjustment ( in this example : Year)

or leave the menu without a change by actuating the CANCEL button .

After the input push the ENTER-button to adopt the changes.

- CONFIG  change or set „Windows“ settings

Warning! Modifications will be overwritten after  Restart with  Standard-values.

To set the modifications as default you have to start the programme „Rflash.exe“ once before powering down. The programme saves the changes in the flash-memory  automatically.

To use this programme, you have to leave the weighing-programmes POWER-ON  by pushing the CANCEL button.

To leave this side  :  push „file“  > „exit“.

4. Language	  choose language

In the window, the shortcut of the chosen language appears.

By pushing the button the display-language changes.

If a language doesn't exist, an error – message is displayed.

Otherwise you will automatically return to the POWER-ON menu.

5.   BACKUP	    backup or restore weighing references

By pushing the BACKUP-button all settings are copied from the

IPC on an USB-stick.

The copied files can be reimported by the RESTORE- button.

6.   UPDATE SERVER   Update main-programme

Push the button EXIT, and after returning into the POWER-ON-menu again

the button EXIT to start the update.

Plug in the USB - Stick and push the OK button.

During the update, the programme will inform you about the copy-procedure.

If all files were copied faultlessly, the following message appears:

Push the button OK to start the weighing – main - programme.

7.    UPDATE CLIENTS    update working - units  (not for the AKE-Units!)

After pushing the button UPDATE CLIENTS  the programme

will import all necessary files.

and copies them to the working-units..

After the checkup you can restart the machine.

Please switch out the main switching device and on again after 10 seconds.

8.   INIT	 Weighing-machine basic settings

After pushing the button INITIALIZATION the menu-window to set

or change the basic settings appears.

Example: push the button „quantity of Wl-s“ and put-in the amount  of  weighing-lines.

After your entry push the button ENTER to adopt it.

9.  MACHINE  Opens the menu of  the weighing-machine

Push the button MACHINE to open the weighing-menu.

The last used settings will be adopted automatically.

Example: Start of WL-1

To start the WL-1 with the available settings push the window STOP

( which is coloured in red now) in the WL-1 nominal frame.

The pushed red window becomes green and the text changes into ON.

10   COUNTER dropcounter settings

By pushing the button COUNTER the numeral keypad opens.

Insert the number of the wanted drops (0-9999) and push the button ENTER.

The counter will change according to your entry and the coloured window changes from red to green.

If you are logged in as a service-person you can insert a number to 9998 to leave the programme and change window settings.

11    OPERATING MODE    set and change operating mode

By pushing the button MODE the menu for the operating mode opens.

Only the pads which are adoptable for your weighing-system can be activated.

- Single-line weighing-machine:                    standard and multiple drops

- Double to fourfold weighing-machine:	standard modes

total drop

total weigh modes

optional: double or multifold drop

If you selected the multifold drop you have to insert the number of drops.

Push the window Drop-Nr. and insert the number of drops in the appeared keypad.

Insert the declaration of weight  as usual  with the button  NOMINAL as if there

was only one drop. The controller will calculate the datas for each drop.

12.    RECIPES     save , import , copy , change , delete recipe

In the recipe-collection the weighing-values are allocated to a name and a number.

That makes the detection over product-name (20 letters) and/or product-number

(8- digit) possible .

All products which name starts with  PROTECT are example databases

which cannot be deleted by  the operator or the foreman.

A copy of this dates is possible without restrictions.

The input of the name will be infected over a alphanumeral keypad

which appears automatically if you  enter , copy or change the name .

If needed , the numeral keypad will be offered automatically.

A selection can be effected by a push of the chosen entry.

over UP and DOWN  or about the SEARCH function.

A product is selected if its name has a blue background.

A grey background  is not an available mark.

By pushing ENTER  the weighing-values for the blue-marked product

are imported  into the memory.

The memory contains 100 entrys maximum in the standard version.

12.a     RECIPE COPY    copy the recipe

To copy an existing recipe you have to select it at first. Select the

recipe by pushing and  push the button COPY .

You will be asked if you really want to copy . Verify the question with  ENTER

if you are sure.  After the verification the alphanumeral keypad appears ,

so you can insert the name of the new entry.

Enter the name of the new entry and verify with ENTER.

To insert the new  product number ( in this example 00000009) and verify

with ENTER.

Enter the new 8-digit product-number ( in this case 00000009) and verify with ENTER.

! Warning ! Every product-number may just be inserted once. If the entered product number is already existing  a message will appear and the procedure will be aborted.

12.b  RECIPE - EDIT    change product-number/-name

1. Change product-name :

After pushing the button EDIT select PRODUCT NAME.

Change or enter the name  and verify with ENTER.

The change will immediatelly appear on the recipe-administrations main-window.

12.c Change product-number.

After pushing the button EDIT choose PRODUCT-NUMBER.

Insert the product number or change it and verify with ENTER:

The change will immediatelly appear on the recipe-administrations main-window.

12.d PRODUKT – DELETE    	delete recipe including ist datas

Delete recipe : Mark the selected entry and push DEL. Verify the

security-question with ENTER or push CANCEL if you want to

stop the deleting- operation.

The deleted entry escapes from the recipe-administrations main window.

! Warning ! The deleting operation is irrevocable . Before you verify the deleting operation you should control if the blue-marked selection really is the right one.

To prevent this error , you should use the backup-function as often as possible.

12.e	RECIPE-SEARCH

1. Searching a product-number

Push the button SEARCH.

Choose PROD.-NR. and put in the searched product-number.

After pushing the ENTER-button the searched product appears blue-marked . If the product does not exist , you will get an error-message.

12.f  Search for product-number

Push SEARCH and choose PROD.-NO.

Insert the name of the searched product.

The searched product appears blue-marked . If the product does not exist ,

you will get an error-message.

13.  TARE-VALUES

Push the button TARE to go to this menu-window.

Taration-interval :  	Number of drops after which an auto-taration will start.

Taration-precision: 	The number between 0-10 regulates the taration´s accurancy .

A smaller number has a higher accuracy as consequence.

Max. Tara Offset:	is not used at the AKE-unit.

Free fill-time:       	At the beginning of the dosing , the free-dosing time

sets the time of the uncontrolled dosage.

This function effects prohibit failures induced by vibrations

on the weighing-bowl..

The Option(s)   	are not used at the AKE-unit.

14. TIMES

Push the button TIMES to go to this menu-window

Start-delay:            Time between the shutters closure on the weighing bowl to the

start  of supply .Prevents dosage into a still open scale.

Measur.-delay      : Time between the supply-shutters closure to

the controlling measure of the amount of product

in the weighing-bowl . This delay makes sure that also

later-arriving products are seized  in the measure.

Threwing off time: Opening-time of the weighing-bowl

Buffer trap delay  :Falling-time into the inter-scale

Dropped on time:   Not used at the AKE-unit

Panning-time:         Not used at the AKE-unit

15.  SUPPLY

Push the button  SUPPLY to go to this menu-window.

Dosing-time supply:     Minimal duration of dosing

Supply to level:	  Value between 0-100% which starts a dosing.

Supply speed:               Not used at the AKE-unit

Panning-time:         	  Not used at the AKE-unit

Option(s):                     Not used at the AKE-unit

16. REFERENCE VALUES

Push the button REFERENCES to go to the menu-window .

Nominal weigh:          The products final weigh

Fine weight:                Weighing value that starts to stop the dosage

All supply-shutters are closing.

Coarse  weigh:            Weighing-value that stops the coarse-weight

and starts the fine-weight.

Pre-dosing value:         Some machines support a fast pre-dosage.

This is only necessary for weighing big capacities.

Upper limit :	              Weighing-value that induces a deflation of the weighing over

a overweight-switch.

Under limit::                (minimum final weigh)

If the final weigh is not reached after the controlling measure ,

an automatical dosage starts until the reference

weigh is reached.

17.  TARE

Push the button TARE to activate this function .The status of the weighing line

intends the following activity.

- If the weighing-system is switched out, a manual taration is necessary.

After a short moment of waiting you get a message.

- If the marked weighing-line is active , only the tarationcounter is set new.

- EMTY

Push the button EMPTY  to activate this function.

Choose between empty weighing-bowl or empty weighing-lines .

You can stop the deflation by pushing the CANCEL button.

19        ON/OFF   weighing-lines start/stop

Push one of the red windows to start a weighing-line .

# The marked window changes into a green colour and ON

WL-1 weighing is ready . Finished-message and real value (including mean

value) have changed to the actual value.

- ACTIVITY ADVERTISEMENT

After a dropping the activity advertisement will be actualized. It shows the average of the last 100 drops.

The inferior windows show the mean values of the last 100 drops.

To delete the mean values (restart) push the according window once.

After pushing, the actual reference value until the first drop will be registerated.

21 CONTROLL controlling-menu

Push the button to go to this menu-window .

With the button TOTAL/NET you can choose between total or net display.

The corner frequency of the digital filter can linear be set between 1-50 hz

by pushing the window FILTER FR.

By pushing the button TARE a chosen weighing-line will be tared .

By pushing the button JUSTAGE a chosen weighing-line an be

justified automatically. ( at the AKE-series the function is limited

to +- 10% correction !

- Tare the justified weighing-line.

- Inscribe the calibrated weight into the window CALIBRATED WEIGHT.

- Push the button JUSTAGE

- Answer to the security question „Has the line already been tared?“

- Put the justifying-weight into the scale.

- Push the weighing-line window.

- Controll as a precaution with the controll-functon if the justification is ok.

8.   If you dealt amiss in this procedure , this mistake can implicate a total error.

Justification-error:

Push the relevant window to reset the multiplier of the display manual .

Insert the multiplier ( given by the company) over the virtual keypad , certify the input with ENTER and repeat the justificaton procedure.

As a correction , a new insert of the multiplier is also possible .To activate this function you have to be logged in as SERVICE-man.

22. INPUT-OUTPUT CONTROL

Push the button PORT in the controlling-menu to go this menu-site.

Outputs                                Inputs

By activating a radio-control the according output is switched on or out.

By pushing the button TEST the outputs will be switched on and out one after another. .

The right key panel shows the status of the inputs.

23.  Optionen           Additional-functions (Optimizing, additional-functions)

optiomation  see:  Description of the optimizing

standard settings are when all functions are radio button off

option:	-1	„on“    ----   overweigh seperator on

-2	„on“    ----   course flap with fine

-3	„on“    ----   initiators deactivated

-4	„on“    ----  predos. By freedosage

-5	„on“    ----   finevibr.s with coarse  value

-6	„on“    ----   finevibr.s with V2, V5 value

-7	„on“    ----   V0 switching to finevalue

-8 	„on“    ----   weightlines with 1 vibrator

- !  the selected options are saved with the other productvalues

22     	Calibrating the touch panel:

a.  Push the button CONFIG on the activating- menu site.

b.  Push the button STYLUS.

- Push the button CALIBRATION

d.  Push CALIBRATE

Push the cross-symbol until it moves.Outgoing from the center of the display

the displays corners are scanned.

After scanning all measuring points , the cross-symbol dissappears and

the following message appears.

With a double-click you can leave this site .

Also leave the CONFIG-menu site and close the main programme

by pushing CANCEL on the activating-menu site.

To save the settings in the flash , open the programme „Flush.exe“

Weigh initialisation

- Change the settings on the SWITCH-ON INIT menu-site.

- Set the multiplier to MAIN MENU-CONTROL site

- Save the settings with the button SAVE on the main menu site.

- Put in all necessary reference values or copy a PROTECT-recipe.

Description of the optimizing  	(validity: 27.08.1995)

Limits		("d" is the multiplier in the weigher)

min. nom. poids:			200d

coarse poids:			128d

fine hyterese:	 		  64d

poids control step upstairs:	1/64 of fine coarse different

downstairs:				1/128 of the coarse poids

vibrator-control step	:		1 position ca. 0,5%

min. dosetime	:		1.2 sec

exercise:	adjust the fundamental adjustements of the weigher so that changments on gutterfillheight, gutteramplitude, and productproperties does not influence the finished poids.

The optimizing is no automatic initialisation or startup automatic!

It is suitable to correct the settings during the working.

operation:   with the menupage which is reachable with key „Q“, at the time 2 modes

are activable

a. coarse vibrator optimization with finevibrator and

coarsepoids optimization

b. coarse vibrator optimization with coarsepoids optimization

but without finevibrator optimization

Before the optimizaton is aktivated the weigher must be setup

by hand and  when power and exactnes are OK,  the working

values may saved with the function "Save status as Reference"

to make them to nom. values for the otimization

On the menupage, reachable with windows of the nom.dosetime

are controlable or changeable by smal steps.

work:	the dosespeeds of the doseunits will be controled.

the coarse dosespeed will be controled with +/-3% precission and
		the fine dosespeed will be controled with +/- 5% precission

coarse dose	fine dose

shorter -3%	shorter -5%	-> 	coarse vibration decreases

shorter -3%	in the range	->	coarse vibration decreases

shorter -3%	longer +5%	->	coarse poids increases

-----------------------------------------------------------------------------------------------

in the range	shorter -5%	->	a. fine vibration decreases

b. coarse poids decreases

in the range	in the range	->	fine poids are corrected

in the range	longer +5%	->	a. fine vibration increases

b. coarse poids increases

----------------------------------------------------------------------------------------------

longer +3%	shorter -5%	->	coarse poids decreases

longer +3%	in the range	->	coarse vibration increases

longer +3%	longer +5%	->	coarse vibration increases

Fine poids correct:

Acttual poids

fewer then nom. Poids       -> 	fine poids = fine poids + (nom.poids – act.poids)/2

more than nom. poids	->	fine poids = fine poids - (act.poids –

nom.poids)/2

in range +/- 1 digit		->	no adjusting

remember:

When big changes happens on the convey to product (silo empty, matarial blockage etc.)
switch the optimizing immediately off, because it trys to correct the errors with too
big adjustements.

Stop the weight lines only in emegency cases because the folowed step after switching on is corrupt ! If desired switch the weight line with the switch off function „A“ off. In this case
the weightline stops afte finished the dose.

Change any adjusting values (nominal poids, vobrator adjustments) only while the optimizing is switched off!

When silo is empty or not completly full (example empty working) switch the optimizing off.

Do not tuch the weightheads while optimizing runs! (in other cases also not recommendable

The productmemory saves the nominal values for optimizing but

switches them not on
	at a product changing. This must be done manualy after starting and

opening the silo!

tip:	when the optimizing function has changed the nom. values with wrong values 
(expl. convey to product disturbs) so you can react the old values with a productchange.